## Slide 1: SiPM読み出しによる 高時間分解能 シンチレーションカウンターの開発
2013年9月20日
日本物理学会2013年秋季大会＠高知大学
素粒子物理国際研究センター
内山 雄祐
西村 美紀, 大谷 航
Matteo De Gerone
他, MEG collaborators

## Slide 2: Motivation
Search for lepton-flavor-violating μ→eγ decay
  Forbidden in SM
  But enhanced in many BSM
MEG experiment
  Searching for μ→eγ down to O(10-13)
  Completed data-taking Aug.2013
  Analysis for final result ongoing
Upgrade
  Push down to O(10-14)
  Approved by PSI, R&D progress
  To start DAQ in 2016
Required to suppress accidental BG
22pSE5 岩本
52.8MeV (＝Mμ/2)
52.8MeV(＝Mμ/2)
Signature

## Slide 3: MEGアップグレード
MEG Upgrade Proposal
(http://arxiv.org/abs/arXiv:1301.7225)
全体像

## Slide 4: SiPM in MEG upgrade
液体キセノンガンマ線検出器
陽電子時間測定器
ガンマ線入射面をPMTから
MPPCに置き換える
真空紫外光に有感
大型
電荷測定・時間測定
～4000ch
高速プラスチックシンチレータ読み出し
近紫外光
高精度時間測定
～500 counters × 2ch (3000 SiPMs)
BG同定検出器
新たに導入を検討
GSO+プラシン ＋ファイバー

## Slide 5: Previous results
  Achieved excellent resolution of 43 ps() at relatively high over-voltage
  Observed deterioration at further over-voltage
    Increase of dark noise, after pulsing
BC422
6x3x0.5 cm3
3M mirror film
Saint-Gobain BC422
(60×30×5 mm3)
SiPM

## Slide 6: Test setup
Scintillator：BC422, 60x30x5mm3
Sensor：3x3mm2, 50μm-pixel SiPMs
Reference counter：BC422, 5x5x5mm3, 1 MPPC for trigger & collimation
90Sr β-source (<2.28MeV)
Amplification & shaping (pole-zero cancellation)
High speed sampling with waveform digitizer
Thermal chamber
Waveform digitizer
Shaping amplifier
3 SiPM in series connection
DRS
USB→PC
Input (4ch)

## Slide 7: Topics today
Further study to understand the timing resolution of plastic scintillation counter from SiPM properties
Detail comparison of different type of SiPMs
  Recently, many manufacturers produce blue-sensitive SiPMs with ‘p-on-n’ structure
  New MPPCs from Hamamatsu
    New standard MPPC
    New trench MPPC
Pursue ultimate timing performance
of general scintillation counter
HPK has recently developed new technologies
After pulse suppression
Trench for cross-talk suppression
Metal quench resistor
High fill factor for small pixel
Etc.
IEEE NSS/MIC 2012,
VCI 2013 Id:180
T.Nagano et.al, IEEE NSS/MIC 2012

## Slide 8: Single sensor properties

## Slide 9: Test samples
Manufacturer | Model number | Type | 
HPK | S10362-33-050C | Conventional MPPC (Old) | Ceramic package
 | S10931-050P |  | Surface mount
 | S12572-050C(X) | New MPPC (Standard) | 
 | S12572-025C(X) |  | 25μm pixel
 | S12652-050C(X) | Trench-type MPPC | 
 | 3X3MM50UMLCT-B |  | Improved fill factor
AdvanSiD |  | NUV type | 
KETEK | PM3350 prototype A | Trench type | 
SensL | MicroFB-30050-SMT | B-series with fast output | 
 | MicroFB-30035-SMT |  | 35μm pixel
Common features:
3×3 mm2 dimension, p-on-n structure

## Slide 10: I-V characteristics
2 V
4 V
＠23℃
Much wider operational range.
Breakdown voltages are lower by 5 V (standard), 15V (trench)
8 V
Trench MPPC
New MPPC
50μm
25μm
Old MPPC

## Slide 11: Snapshot of waveform
Confirmed less after-pulsing
New MPPC (S12572-050C(X))
  Old MPPC (S10362-33-050C)
Average waveform
20 ns
After shaping
Trench MPPC (S12652-050C(X))

## Slide 12: p.e. counting
Dark measurement
  Random trigger
Photo-electron counting capability
  Large improvement in the separation
0pe
1pe
2pe
Over-voltage ~2.0 V
Over-voltage ~2.5 V
Charge (109e)
アフターパルスにより連続的な分布に
Old MPPC
New MPPC

## Slide 13: Dark count
New
Old
＠23℃
New MPPCs
Old MPPCs
Kept controlled up to high Vover due to the after-pulse suppression
SensL
AdvanSiD
KETEK

## Slide 14: Cross talk
＠23℃
Cross-talk
Probability
Standard SiPMs
Trench SiPMs

## Slide 15: Cross talk
Cross talk increases at high bias
  Very long tail in charge distribution
  Cross-talk process runaway
  Big impact on energy measurement (excess noise factor)
  Impact on time measurement?
Vover=1.5V
Vover=4.0V
1pe
2pe
…
10pe
100pe
1pe
5pe
Charge (109e)
Saturation of electronics
2pe
Standard NEW MPPC
arXiv:1109.2014

## Slide 16: PDE with UV-LED
Measure relative PDE using UV-LED (~ scinti. Light)
Calculate PDE from 0 p.e. probability
LED (~390nm)
LED2 (~365nm)
Trench
25um
New50um
Old
SensL
KETEK
AdvanSiD
Extended PDE at higher Vover for new MPPCs
Recovered fill factor for smaller pixel (25um)
Trench-type MPPC also shows good PDE due to relatively high fill factor (55%, ~10% lower than standard type)
Scintillator spectrum
(Photo Detection Efficiency)

## Slide 17: PDE with UV-LED
Measure relative PDE using UV-LED (~ scinti. Light)
Calculate PDE from 0 p.e. probability
LED (~365nm)
LED2 (~365nm)
Trench
25um
New50um
Old
SensL
KETEK
AdvanSiD
Extended PDE at higher Vover for new MPPCs
Recovered fill factor for smaller pixel (25um)
Trench-type MPPC also shows lower PDE for short wavelength (?)
Scintillator spectrum

## Slide 18: PDE with UV-LED
LED2 (~365nm)
LED (~365nm)
Trench
25um
New50um
Old
SensL
KETEK
AdvanSiD
HPK MPPCs show higher PDE for near-UV
Extended PDE at higher Vover for new MPPCs
Recovered fill factor for smaller pixel (25um)
Trench-type MPPC also shows good PDE due to relatively high fill factor (55%, ~10% lower than standard type). But PDE is suppressed at shorter wave length (?)
LED (~390nm)
Trench
25um
New50um
Old
SensL
KETEK
AdvanSiD
*Relative scales for the two plots are not same

## Slide 19: Scintillator signal readout by 3 MPPC series connection

## Slide 20: 3 series connection
直列につないだ3つのセンサーを1チャンネルで読み出す
  実効的に大面積化。検出光量の増加で分解能をあげる。
  両端に3倍のバイアス電圧を供給。
共通電流で個々のセンサーにかかる電圧が調整される
  Over-voltageが自動的に大体そろう
直列接続でキャパシタンス減少
  シャープな波形
  ゲイン減少 (~40%)
Over-voltage
~3.8 V /MPPC

## Slide 21: Time resolution
New (25μm)
Almost same resolution attained at each optimal bias. (No major improvement)
Trench-type shows slightly poorer resolution
New
Old
Trench
BC422
6x3x0.5 cm3 No wrapping

## Slide 22: Rise time
New (25μm)
New
Old
Trench
Observe softening of waveform as Vover for standard new MPPC
Slower pulse shape for Trench MPPC
  This sample is not the final version.
  HPK is trying to improve this

## Slide 23: Pulse timing
New (25μm)
New
Old
Trench
Faster charge collection by higher E-field
As a consequence of softened pulse shape, time measurement drifts for applied bias
→ Larger temperature dependence of time measurement
!？
Pulse time by Constant-Fraction method
Old | 2.5 ps/℃
New | 5.5 ps/℃
New 25um | 2.8 ps/℃
Trench | 0.1 ps/℃

## Slide 24: Temp vs Timing
20°C
25°C
30°C
20°C
40°C
30°C
New MPPC (S12572-050C(X))
  Old MPPC (S10362-33-050C)
バイアス一定の場合の時間測定(mean)の温度依存性
高バイアス下での時間のドリフト(波形のなまり)により新型では温度依存性が大きくなってしまった。(5度の変化で~30ps)
Bias voltage (V) (Over voltage/MPPC @23°C)
Bias voltage (V) (Over voltage/MPPC @23°C)
30 ⊕ 30/√12 = 31.2 ps
検出器の分解能(最終目標値)
温度変化の寄与(一様分布仮定)
問題となるレベルではない
Temp. dependence of time measurement (mean) at fixed bias voltage
Due to the time drift (pulse shape softening) new MPPC has larger dependence (~30ps by 5℃ variation)

## Slide 25: AdvanSiD
20°C
25°C
30°C
20°C
25°C
30°C
  Breakdown coefficient
  AdvanSiD: 24 mV/℃
  KETEK： 16 mV/℃
Bias voltage (V) (Over voltage/MPPC @23°C)
Bias voltage (V) (Over voltage/MPPC @23°C)
MPPC attained
Stable!
  MPPC Old: 49 mV/℃
  MPPC New:59 mV/℃
10–15 ps poorer resolution
Good stability for wide bias range
  No more care for temperature

## Slide 26: まとめ
New MPPC
  after pulseの大幅削減, 動作領域の拡大
    パルス後の安定性向上。高計数率下測定で有効。
  Trench-type, 加えてcross talk抑制
  Small pixelでもfill factor改善で同等のPDE
時間測定の温度依存性
  New MPPCで悪化
  Trench-typeではKETEK, AdvanSiD同様安定
シンチレータ時間分解能
  p.e. statisticsがドミナント → PDEが重要
    近紫外光で高いPDEを有する浜松MPPCで高分解能
  New MPPCでは到達分解能に大幅な改善は見られない。

## Slide 27
 | After pulse, noise | Cross talk | PDE | Bias, Temp. stability | Time resolution
Old MPPC | △ | ▼ | ○ | △ | ◎
New MPPC (Standard) | ◎ | ▼ | ◎ | ▼ | ◎
New MPPC (25um) | ◎ | ▼ | ◎ | △ | ◎
New Trench MPPC | ◎ | ◎ | ○ | ◎ | ○
AdvanSiD NUV | ▼ | ○ | ▼ | ◎ | △
KETEK | ▼ | ◎ | ▼ | ◎ | ▼
時間分解能追求ならStandard MPPC (50 or 25 μm pixel)
総合性能ならTrench MPPC
◎ ○ △ ▼
Good

## Slide 28: Scintillator
Faster scintillator gives higher time resolution
  P.e. at the earliest part are only effective
  Scintillation with fast response given in near-UV light
properties | BC418 | BC420 | BC422 | BC404
Light Output [% Anthracene] | 67 | 64 | 55 | 68
Rise Time [ns] | 0.5 | 0.5 | 0.35 | 0.7
Decay Time [ns] | 1.4 | 1.5 | 1.6 | 1.8
Wavelength [nm] | 391 | 391 | 370 | 408
Attenuation Length [cm] | 100 | 110 | 8 | 140
Present MEG timing counter
Used in the beam test
Used in this study (Best )
Fast scintillators from Saint-Gobain

## Slide 29: 1 pe signal
New Standard, 4V
New Trench, 4V
New Trench, 9V
Trench型で大幅な波形のなまりを観測。

## Slide 30: Temp. vs Resolution
バイアス一定の場合の分解能の温度依存性
  レファレンスカウンタは各温度でOver-voltageをそろえてある
新型では高バイアスをかけることで、到達分解能は安定に得られる。
20°C
25°C
30°C
20°C
40°C
30°C
New MPPC (S12572-050C(X))
  Old MPPC (S10362-33-050C)
Bias voltage (V) (Over voltage/MPPC @23°C)
Bias voltage (V) (Over voltage/MPPC @23°C)

## Slide 31: Temp. vs Resolution
Over voltage一定での分解能温度依存性。
  新型(旧型)60mV/°C(55mV/°C)の温度係数を仮定して補正。
20°C
25°C
30°C
20°C
40°C
30°C
New MPPC (S12572-050C(X))
  Old MPPC (S10362-33-050C)

## Slide 32: 温度依存性・時間測定
Bias voltage (Over voltage/MPPC @23°C) (V)
20°C
25°C
30°C
20°C
40°C
30°C
New MPPC (S12572-050C(X))
  Old MPPC (S10362-33-050C)
Bias voltage (Over voltage/MPPC @23°C) (V)
Over voltage一定での時間測定温度依存性。
  新型(旧型)60mV/°C(55mV/°C)の温度係数を仮定して補正。
もし温度変化をバイアスコントロールでアクティブに補正できれば影響を最小限にできる。

## Slide 33: 温度依存性
降伏電圧の温度依存性
  Old:49 mV/℃
  New:59 mV/℃
New MPPC (S12572-050C(X))
  Old MPPC (S10362-33-050C)
参考:
AdvanSiDセンサー:24 mV/℃
KETEKセンサー：16 mV/℃
(どちらもp-on-n, blue sensitive type)

## Slide 34: 整形・増幅器
入力段減衰
１５０
ポールゼロ消去
MPPCバイアス電圧
MPPC
アンプ
アンプ
出力

## Slide 35: Shaping
For precise time measurement
  Restore baseline by pole-zero cancellation, and
  Extract fast rise-part
200 Ω
Pole-zero cancellation potentiometer 0 Ω (attenuated by 2)
1000 Ω
0 (nsec)
-50
-100
-150
trise~1.7 ns
Signal from scintillator （３MPPC in series connection）

## Slide 36: Excess noise factor
Old
Vover=2.4V
Pcrosstalk=60%
New
Vover=2.5V
Pcrosstalk=50%
New
Vover=4.2V
Pcrosstalk=70%
Trench
Vover=8V
Pcrosstalk=35%

## Slide 37: Gain
数光子ピークからゲインの見積もり
  バイアス依存性
    線形性はかねがねよい。
    高バイアスでずれ？
  同じover-voltageではゲインは従来型より低い。
New
Old
Total gain includes amp, splitter, offline shaping.
＠23℃