## Slide 1: Mass test of MPPC with a prototype of MEG II liquid xenon detector
家城 佳, 他 MEG-II collaboration
II

## Slide 2: MEGII liquid Xe (LXe) detector
2inch PMT× 216
12x12mm2 MPPC × ~4000
1m
inner face
upgrade!
(CG)
Granularity will improve
 γ energy resolution & position resolution expected to improve by a factor of 2!
Requirements for the MPPC:
- sensitive to VUV scintillation light
- large area (less number of readout ch)

## Slide 3: VUV sensitive large area MPPC
Successfully developed a new type of MPPC! - PDE for VUV light > 20% - Large area
fall time < 50ns with series connection
- 50μm pitch pixel
Hamamatsu S10943-3186(X)
- metal quench resister
- Four chips
Performance tests so far:
- Mass test (600 pcs) in room temperature
- Test of small subsamples in LXe
This talk: mass test (~600 pcs) in LXe

## Slide 4: Mass test in LXe
α sources
241Am
LED
New PCB and new feed-through
feed-through
PCBs
(12+12)
12mm
MPPC x 568 pcs
pre-final model
(no crosstalk suppression)
We would like to check if everything work OK in LXe.
PCB
24MPPCs / PCB
two types of series connection of four chips for each MPPC
In LXe
2seg
4seg
or

## Slide 5: Setup
MPPC &
support structure
feed-through
Xe gas panel

## Slide 6: Setup
jigs to mount / dismount MPPCs
feed-through
MMCX cable
cable remover
MPPCs mounted on PCBs
It was difficult to handle small connectors in a limited space.
3mm

## Slide 7: Data taking
LED trigger
Vover ~ 2.6-3.3 V, 0.1V step
Alpha trigger
Vover ~ 3.0V
 Breakdown voltage (Vbd), gain, crosstalk+afterpulse (CTAP)
 PDE
Waveforms are recorded with waveform digitizer.
24 channels are read out for each measurement.
We finished the measurement on 7 Mar.  Preliminary results in this talk
Charge
100ns
5mV
Example of waveform
charge integration
Example of charge distribution
0p.e.
1p.e.
….

## Slide 8: Results: breakdown voltage
Looks OK except for some bad channels.
I will explain this later
Example of 1p.e. charge vs. Vbias
Vbd vs. serial no.
Vbd (measured)
Vhama – 10.5V
Preliminary

## Slide 9: Results: gain
Looks OK except for bad/strange channels.
Preliminary
4seg serial connection
2seg serial connection
Gain for all MPPCs
@ Vover = 3.0V

## Slide 10: Results: crosstalk + afterpulse
Looks OK except for bad/strange ch.
CTAP prob. for all MPPCs
1p.e.
2p.e.
3p.e.
CTAP effect
CTAP prob.
=
Expected num of 1p.e. events
Measured num of 1p.e. events
Estimated from 0p.e. peak, assuming Poisson statistics.
Preliminary
1 -
@ Vover = 3.0V
0p.e.
Charge

## Slide 11: Results: PDE
PDE = (Measured p.e.) / (Expected p.e.)
peak of charge distribution
corrected for CTAP
Very Preliminary
The measured PDE is lower than the PDE which we measured in the small chamber (~20%).  We will check this in more detail.
α energy / Xe scintillation yield * solid angle
Example of alpha charge distribution
Distribution of PDE
@ Vover ~ 3.0V

## Slide 12: Bad channels
Several types of bad channels are found
good channel
dead channel
strange channel
x 16
x 9
7 out of 16 dead channels were OK in the signal check which we did before liquefaction of Xe.
We found 8 dead channels in the signal check before liquefaction. 5 of them look OK now.
It is likely that the connections of MMCX cables are bad at the PCBs or at the feed-through. We will check this in more detail.
We also plan to improve the design of the PCB and feed-through.

## Slide 13: Summary
We performed a mass test of MPPCs in LXe. New PCB and feed-through were also tested.
Preliminary results of the test is presented. Performance of most of the channels looks OK, as we expected from the previous tests.
We found some bad channels. It is likely that the connections of the cables are bad. We will check this in more detail. Design of the feed-through will be improved.

## Slide 14: Backup

## Slide 15: PCB
3列2列に変更

## Slide 16: Feed-through
最初からケーブルが半田付けされた
デザインに変更する。

## Slide 17: PDE vs. distance
αのほぼ真下のchの光量分布
直近のαのピーク
(距離~6cm)
離れたα wireのピーク
(距離~12cm)