## Slide 1: MEG II実験液体キセノン検出器の インストール及びセンサー試験
家城 佳、
他 MEG II コラボレーション
 次のトークで

## Slide 2: search
MEG
MEG II
DeeMe
COMET-I
Mu3e-I
Mu3e-II
COMET-II
Mu2e
μNeN
LFV upper limit history
Observation of lepton flavor violation (LFV)
decay
Evidence of new physics!
New generation LFV experiments can explore the region predicted by BSM
BSM prediction
e.g. with SUSY

## Slide 3: MEG II
μ+
γ
e+
beam @ PSI
~7x107 μ/s
900 LXe detector
drift chamber + timing counter
Radiative decay counter
(BG identification)
Upgrades from MEG:
x2 beam rate
x2 detector resolution and efficiency

## Slide 4: LXe detector
Measures the of -rays.
Readout: 2” PMTs 668 + 12mm SiPMs 4092  Granular& uniform coverage  Resolution will be improved from MEG
timing
position
energy
1 m
2 inches
12 mm
Energy resolution ~1%
Position resolution ~2.5mm
PMTs on this surface are replaced with MPPC

## Slide 5: Final stage of construction
Cabling
Closing the lid
Closing the flanges

## Slide 6: Installation & pipe connection
μ
GM
refrigerator
(new)
Installation
Pipe connection
LXe transfer tube

## Slide 7: Filling LXe to the detector
Liquefaction (GXe storage tanks  1000 LXe tank)
Vacuum evacuation of the detector (to ~2x10-3 Pa)
Pre-cooling of the detector (~175K)
LXe transfer from 1000 tank to the detector
High pres.
GXe tanks
x10
1000l
LXe tank
(cooled by refrigerator & LN2)
purifier
(getter)
detector

## Slide 8: LXe transfer
View from a USB camera installed inside the detector

## Slide 9: Position monitoring
outer cryostat
inner cryostat
(filled with LXe)
MPPC position was scanned with laser after installation↓
However, inner cryostat deforms by heat shrink and LXe load.
 Monitor the movement by position sensors
vacuum
wire position sensors
are attached to the outer cryostat.
wire pull out length is measured

## Slide 10: Position monitoring
pre-
cooling
LXe transfer
3 sensors are installed at 4 different places.
Moved mostly during the pre-cooling due to heat shrink. Result was roughly consistent with what we expect (~1.6x10-3mm/m/K heat shrink to bottom direction)
Cryostat is fixed at the bottom middle.
observed movement
Movement (top US sensor)

## Slide 11: X-ray survey
57Co source + collimator
X-ray
(124, 132 keV)
moving stage & rotator
Another method for position check: Measure MPPC signal from the X-rays.
Position of the stage is measured by laser and monitored by laser and bubble level + camera.

## Slide 12: X-ray survey
15mm
Example of X-ray signal at one MPPC
Scan was performed in two directions in 1mm step.
X-ray spot size ~ 2 mm x 30 mm
scanned region
Background (cosmic-rays)
Event rate increase was successfully observed around X-ray irradiated region.
Analysis is ongoing.
X-ray beam
MPPCs
15mm

## Slide 13: LXe monitoring & control
SCS2000
We built a slow control system using special modules (SCS2000) developed at PSI
GM refrigerator
pulse-tube refrigerator
pump
purifier
(getter)
pres. sensors
temp. sensors
readout
feedback
SCS2000 can:
- readout the sensors
- automatically or manually control the cooling power (refrigerators and LN2)
- allow users to monitor & control via internet

## Slide 14: LXe monitoring & control
refrigerator
electric
heater
LN2 pipes
Normal (stable) operation: Temperature & pressure is maintained by adjusting electric heater power.
Emergency or special operation: Open LN2 valves for additional cooling
Firmware of SCS2000 does the control automatically.
- Heater power adjustment (PI control)
- LN2 valve open at high temp. or high pres.
- Issue an alarm in case of emergency

## Slide 15: LXe monitoring & control
Overall system is monitored & controlled from web browser.

## Slide 16: Purification
pump
getter
GM refrigerator
pump
molecular sieve
pulse tube refrigerator
Gas phase purification
Liquid phase purification
Gas phase (getter): Slow, remove H2O, O2, CO, CO2, N2, H2 … Liquid phase (molecular sieve):
Fast, remove H2O
Two types of purifications are ongoing in parallel.
Purity of Xe affects the scintillation (light yield, attenuation, pulse shape)

## Slide 17: Detector stability
Xe purity is improving (light yield increased ~10 times over 2 month.) Currently the light yield is ~85% of what we achieved in MEG.
Temperature and pressure are sometimes unstable due to LXe purification pump trouble and due to heat income in LXe purification line.
 Once we finish purification, the detector will be more stable.
Pressure variation (1 week)
Temperature variation (1 week)
pump trouble
0.001MPa
0.04 deg

## Slide 18: Status & schedule
construction
commissioning & purification
engineering run
we are here
Construction is finished!
LXe operation started. Purification is ongoing.
MPPC & PMTs are being tested  Next talk (S.Ogawa) Only a limited amount of DAQ channels is available now (~1000ch)
Detector monitoring is started  Next talk (N. Matsuzawa)
Engineering run will start in the end of 2017
2017
2018
physics run

## Slide 19: Summary
LXe detector is successfully installed.
Position of the detector is measured by wire sensors and X-ray.
Slow control system is built. Purification is ongoing. (so far achieved ~85% of light yield in MEG)
Commissioning of MPPC & PMTs is also ongoing.
Monitoring of the detector is started.