# Xenon Photon Detector

# Cryostat Design Summary

Satoshi Mihara, Tom Haruyama

22 November 2003

Revised January 2004

## Introduction

This note describes summary of cryostat design for the photon detector done by Japanese group in collaboration with a design/construction company, SAAN in Japan. Information described in this note can be found also at the project web site http://meg.web.psi.ch/subprojects/calorimeter/calorimeter.html.

## Overview of the cryostat

The cryostat consists of two layers, outer vessel and inner vessel. The inner vessel is supported from the outer vessel by G10 supports so that heat conduction can be reduced as small as possible. There are three large nozzles on the top of the cryostat where equipments can be mounted like a vacuum pump, feed-through connectors for temperature readout; HV supply and PMT signal readout, pressure/vacuum gauges, and various kinds of valves for xenon transfer. A pulse-tube refrigerator will be mounted on one of the nozzle for cooling. A liquid nitrogen-cooling pipe is planned to be attached on the cover as done in the large prototype,. An overview of the cryostat can be found in Fig. 1.

Fig. 1  Overview of the cryostat. The drawing can be found at　 http://meg.web.psi.ch/subprojects/calorimeter/drawings/210overview.dwg

## Important Parameters

### Strength Calculation for the walls

Design values of wall thickness were determined based on the following calculations. Details are shown in separate tables from Table 1 to Table 4 and are summarized in Table 5.

There are in principle 12 walls on the cryostat,

- Inner vessel

  - Inner wall, Outer wall (cylindrical)

  - 2 Covers at both ends (plate)

  - 2 front walls at top and bottom of the cryostat (plate)

- Outer Vessel

  - Inner wall, Outer wall (cylindrical)

  - 2 Covers at both ends (plate)

  - 2 front walls at top and bottom of the cryostat (plate)

The minimum thickness of each wall was calculated under the condition of vacuum in the insulation layer and 0.3MPa in the inner vessel. Detector operation will be done in principle around 0.1MPa. During liquefaction of xenon the pressure can reach 0.2MPa but never above that except for emergency like refrigerator operation failure due to power cut (even in this case we have liquid nitrogen cooling system for avoiding the disaster). Thus 0.3MPa design still has some margin of pressure tolerance.

| Cylindrical Body Wall |  |  |
| --- | --- | --- |
|  | Formula | t=P*Di/(2*_a* -1.2*P) |
|  |  | Outer Vessel Inner Wall | Inner Vessel Outer Wall |
| Input Parameter | Design Pressure (MPa) | 0.1 | 0.3 |
|  | Design Temp (deg C) | 40 | 40 |
|  | Material | SUS316L | SUS316L |
|  | _a at the design Temp (N/mm2) | 109 | 109 |
|  | Welding efficiency  | 1 | 1 |
|  | Inner Radius of the body Di (mm) | 1100 | 2218 |
| Result | (I) 2*_a* | 218 | 218 |
|  | (II) 1.2*P | 0.12 | 0.36 |
|  | (III) (I)-(II) | 217.88 | 217.64 |
|  | (IV) P*Di | 110 | 665.4 |
|  | T=(IV)/(III) | 0.5 | 3.1 |
|  | Minimum thickness (mm) | 0.5 | 3.1 |
|  | Actual thickness (mm) | 0.5 | 7.0 |

Table 1  Wall thickness calculation of Outer-Vessel Inner wall and Inner-Vessel Outer wall where pressure direction is from the center to the outside.

| Cylindrical Body Wall |  |  |
| --- | --- | --- |
|  | Formula | t1 = 3*P*D0/(4*B*C) for t1>=t0 |
|  |  | t2 = 3*P*D0/(4*A*E*C) for t2>=t0 |
|  |  | Outer Vessel Outer Wall | Inner Vessel Inner Wall |
| Input Parameter | Design Pressure (MPa) | 0.1 | 0.3 |
|  | Design Temp (deg C) | 40 | 40 |
|  | Material | SUS316L | SUS316L |
|  | Breakdown point (N/mm2) | 154 | 154 |
|  | Outer diam of the body D0(mm) | 2400 | 1232 |
|  | Coeff. depending on welding C | 1 | 1 |
|  | Length of the body L(mm) | 1270 | 1086 |
|  | Supposed minimum thickness t0(mm) | 5 | 5 |
| Result | L/D0 | 0.52916667 | 0.88149351 |
|  | D0/t0 | 480 | 246.4 |
|  | Coeff. determined by the shape A | 0.00025 | 0.00052 |
|  | Coeff. depending on the material B | 24 | 51 |
|  | (I) 4*B*C | 96.0 | 204.0 |
|  | (II) 3*P*D0 | 720.0 | 1108.8 |
|  | t1 = (II)/(I) | 7.5 | 5.4 |
|  | t1 >= t0 ? | OK | OK |
|  | Minimum thickness (mm) | 7.5 | 5.4 |
|  | Actual thickness (mm) | 8 | 7 |

Table 2  Wall thickness calculation of Outer-Vessel Outer wall and Inner-Vessel Inner wall where pressure direction is from the outside to the center.

| Plates fixed to a cylindrical body with bolts |  |  |
| --- | --- | --- |
|  | formula | t=G*sqrt(0.3*Z*P/_a + 6*W*hG/_a*L*G^2) |
|  |  | Inner Vessel Cover | Outer Vessel Cover |
| Input Parameters | Design Pressure P(MPa) | 0.3 | 0.1 |
|  | Design Temp T(deg C) | 40 | 40 |
|  | Material | SUS316L | SUS316L |
|  | _a at the design Temp (N/mm2) | 109 | 109 |
|  | G of Gasket G(mm) | 494 | 580 |
|  | Moment Arm hG(mm) | 18 | 20 |
|  | Weight on bolt W(N) | 61054 | 27326 |
|  | Peripheral length of bolt hole L(mm) | 5881 | 6750 |
|  | Diameter or minimal span d (mm) | 560 | 660 |
|  | Maximum span D(mm) | 2365 | 2373 |
| Result | Z=3.4-2.4*G/D, Z<=2.5 | 2.5 | 2.5 |
|  | (I) 0.3*Z*P | 0.225 | 0.075 |
|  | (II) (I)/_a | 0.00206422 | 0.00068807 |
|  | (III) 6*W*hG | 6593832 | 3279120 |
|  | (IV) _a*L*G^2 | 1.5643E+11 | 2.4751E+11 |
|  | (V) (III)/(IV) | 4.2151E-05 | 1.3249E-05 |
|  | (VI) sqrt((II)+(V)) | 0.04589522 | 0.02648249 |
|  | T=G*(VI) | 22.6722377 | 15.3598415 |
|  | Minimum thickness (mm) | 22.6722377 | 15.3598415 |
|  | Actual thickness (mm) | 30 | 20 |

Table 3 Wall thickness calculation for the Inner-Vessel and Outer-Vessel covers.

| Plates |  |  |  |
| --- | --- | --- | --- |
|  | formula | t=d*sqrt(Z*C*P/_a*) |
|  |  | Inner Vessel Front Wall | Outer Vessel Front Wall |
| Input Parameters | Design Pressure P(MPa) | 0.3 | 0.1 |
|  | Design Temp T(deg C) | 40 | 40 |
|  | Material | SUS316L | SUS316L |
|  | _a at the design Temp (N/mm2) | 109 | 109 |
|  | Welding efficiency  | 1 | 1 |
|  | Coeff. depending on attachment C | 0.33 | 0.33 |
|  | Diameter or minimal span d (mm) | 504 | 650 |
|  | Maximum span D(mm) | 1086 | 1270 |
| Result | (I) _a* | 109 | 109 |
|  | (II) Z=3.4-2.4*d/D | 2.28618785 | 2.17165354 |
|  | (III) Z*C*P | 0.2263326 | 0.07166457 |
|  | (IV) sqrt((III)/(I)) | 0.04556804 | 0.02564124 |
|  | t=d*(IV) | 22.9662898 | 16.6668048 |
|  | Minimum thickness (mm) | 22.9662898 | 16.6668048 |
|  | Actual thickness (mm) | 24 | 20 |

Table 4  Wall thickness calculation for the front plates of the inner and outer vessel.

| Walls/Cover | Material | Dimension | Pressure Direction | Design P [MPa] | Design T [degree C] | Thickness Calc [mm] | Thickness Act [mm] |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Inner Vessel Inner Wall | SUS316L | R623 x L1086 | Outer | 0.3 | -108~40 | 5.4 | 7.0 |
| Inner Vessel Outer Wall | SUS316L | R1116 x L1086 | Inner | 0.3 | -108~40 | 3.0 | 7.0 |
| Inner Vessel Front Wall | SUS316L | min504 x max1086 | Inner | 0.3 | -108~40 | 23.0 | 24.0 |
| Inner Vessel Cover | SUS316L | R1148 x R588 | Inner | 0.3 | -108~40 | 22.7 | 24.0 |
| Outer Vessel Inner Wall | SUS316L | R550 x L1260 | Inner | 0.1 | 40.0 | 0.5 | 0.5 |
| Outer Vessel Outer Wall | SUS316L | R1200 x L1270 | Outer | 0.1 | 40.0 | 7.5 | 8.0 |
| Outer Vessel Front Wall | SUS316L | min650 x max1270 | Outer | 0.1 | 40.0 | 16.7 | 20.0 |
| Outer Vessel Cover | SUS316L | R1200 x R540 | Outer | 0.1 | 40.0 | 15.4 | 20.0 |

Table 5 Summary of wall thickness calculations

### Strength calculation for the G10 support

Strength calculation for the G10 support is summarized in Table 6.

| Strength Calculation for the G10 support | Unit | Formula | Value |
| --- | --- | --- | --- |
| W1 | Inner Vessel weight | kg |  | 1200 |
| W2 | Liquid xenon weight | kg |  | 2400 |
| W3 | PMT weight | kg |  | 400 |
| W | total weight | kg | W1+W2+W3 | 4000 |
| n | number of support |  |  | 2 |
| w | weight/support | kg | W/n | 2000 |
| d1 | outer diameter of the support | mm |  | 115 |
| d2 | inner diameter of the support | mm |  | 95 |
| s | allowed strength of G10 | kg/mm2 |  | 10 |
| A | cross section | mm2 | pi/4(d1^2-d2^2) | 3298.672 |
| Wr | allowed weight |  | s*A | 32986.72 |
| Wr-w | never negative | kg | Wr-W | 30986.72 |

Table 6  Strength calculation for the G10 support.

### Estimation of the total weight

Current estimation of the weight is summarized in Table 7.

|  |  | Material | Dimension | Qty | kg | kg |
| --- | --- | --- | --- | --- | --- | --- |
| Inner Vessel | Inner Wall | Honeycomb | 0.94g/cm2 | 1 | 15.0 | 15.0 |
|  | Outer Wall | SUS | 307.8x108.6x0.7cm3 | 1 | 185.0 | 185.0 |
|  | Front Wall | SUS | 107x50.8x2.4cm3 | 2 | 103.0 | 206.0 |
|  | Flange | SUS | see the drawing | 2 | 60.0 | 120.0 |
|  | Cover | SUS | see the drawing | 2 | 265.0 | 530.0 |
| Outer Vessel | Inner Wall (thick) | SUS | 126x18.5x2.0cm3 | 2 | 36.9 | 73.7 |
|  | Inner Wall (thin) | SUS | 127x115x0.05cm3 | 1 | 5.8 | 5.8 |
|  | Outer Wall | SUS | 354x127x0.8cm3 | 1 | 285.0 | 285.0 |
|  | Front Wall | SUS | 126x65x2.0cm3 | 2 | 129.5 | 259.0 |
|  | Flange | SUS | see the drawing | 2 | 111.0 | 222.0 |
|  | Cover | SUS | see the drawing | 2 | 260.0 | 520.0 |
| Legs | Plate | SUS | see the drawing | 1 | 327.7 | 327.7 |
|  | Legs | SUS | see the drawing | 3 | 24.5 | 73.5 |
| Subtotal |  |  |  |  |  | 2822.7 |
| PMT | 500 g/PMT including holder |  | 800 | 0.5 | 400.0 |
| Subtotal |  |  |  |  |  | 400.0 |
| Scintillator |  | xenon |  | 800 | 3.0 | 2384.0 |
| Subtotal |  |  |  |  |  | 2384.0 |
| Total |  |  |  |  |  | 5606.7 |

Table 7  Estimation of the weight of the detector.

### Heat load calculation

Heat load calculation is summarized in Table 8. In this calculation it is supposed that 30 layers of super insulators are placed on the inner vessel. Most of heat load is from PMTs bleeder circuit and conduction through cables. We can expect 150W cooling power of the refrigerator is realistic which is sufficient to compensate this static heat load.

| Heat load calculation for the xenon photon detector vessel |  |  |
| --- | --- | --- |
|  |  |  | Unit |
| Radiation | Outer Vessel -> Inner Vessel (30 Mylar layers) | 3.1 | W |
| Conduction | Nozzle (N1-N3) via xenon gas | 0.2 | W |
| Conduction | Nozzle (N1-N3) via bellows | 4.6 | W |
| Conduction | Support (brace and supporting pipe) | 6.3 | W |
| Heat generation | PMT (65mW/PMT) | 52 | W |
| Conduction | PMT HV and Signal Cables | 50 | W |
| total |  | 116.2 | W |

Table 8  Heat load calculation done based on the current design.

## Bolts for fixing the covers

Number of bolts and their positions for fixing the covers are determined based on JIS B8273:1993 “Bolted flange for pressure vessels”. Location of the groove for the metal seal is also determined based on this standard. The shape for the groove is determined by following recommendation from a possible metal gasket producer USUI. Stress calculation is performed to check that the actual stress caused by moment of force is less than the allowed stress of the material. The results are summarized in Table 9.

| Symbol | Description | Unit | Inner Vessel Flange | Outer Vessel Flange |
| --- | --- | --- | --- | --- |
| P | Design Pressure | ＭＰａ | 0.3 | 0.1 |
| f | Allowable Tensile Stress | N/mm2 | 115 | 115 |
| A | Outer Radius | mm | 2296 | 2400 |
| B | Inner Radius (Outer) | mm | 2204 | 2160 |
| B1 | Inner Radius (Inner) | mm | 1276 | 1190 |
| K | A/B |  | 1.042 | 1.111 |
| Y | Coefficient from A |  | 47.0 | 18.4 |
| g0 | Thickness of the leading edge | mm | 7 | 16 |
| b0 | Basic width | mm | 0.7 | 0 |
| b | Effective width | mm | 0.7 | 0 |
| G | Basic width of gasket（outer） | mm | 2,225 | 2,330 |
| G1 | Basic width of gasket（inner） | mm | 1239 | 1,170 |
| m | Gasket Coefficient |  | 4.5 | 0.0 |
| y | Minimum tightening stress | N/mm2 | 150 | 0 |
| b | Allowable Tensile Stress of Bolts | N/mm2 | 102 | 102 |
| C | Radius of bolt center | mm | 2266.0 | 2370.0 |
|  | Bolt Size |  | M20 | M20 |
| D | Core Diameter | mm | 17.294 | 17.294 |
| N | Number of Bolts |  | 70 | 41 |
| tm | Groove Depth | ｍｍ | 4.5 | 7.0 |
| H | Area of G x Ｐ |  | 377,446 | 153,208 |
| Hp | ｂ(Ｇ + G1)ｍＰ |  | 10,279 | 0 |
| H + Hp |  |  | 387,725 | 153,208 |
| Wm1＝W0 | H + Hp |  | 387,725 | 153,208 |
| ｂＧｙ | ｂ(Ｇ + G1)ｙ/2 |  | 571,040 | 0 |
| Ｗm2 | ｂＧｙ |  | 571,040 | 0 |
| Ａm1 | Ｗm１/b |  | 3,801 | 1,502 |
| Am2 | Ｗm2/ａ |  | 5,598 | 0 |
| Ａm | Larger one between Am1 and Am2 |  | 5,598 | 463 |
| Ａｂ | /4 x ｄ＾2 x ｎ |  | 16,435 | 9,626 |
| Ab-Am | Never negative |  | 10,837 | 9,163 |
| Wg | （Ab + Am）/2 x b |  | 1,123,662 | 514,537 |
| HＤ | Area of B x Ｐ |  | 333,935 | 116,121 |
| HG | Wo - Ｈ |  | 10,279 | 0 |
| HＴ | Ｈ - ＨＤ |  | 43,511 | 37,087 |
| hＤ | （Ｃ - Ｂ）/２ |  | 31 | 105 |
| hＧ | （Ｃ - Ｇ）/２ |  | 20.5 | 20 |
| hＴ | （hＤ + ｈＧ）/２ |  | 25.75 | 62.5 |
| MＤ | ＨD x ｈＤ |  | 10,351,992 | 12,192,718 |
| MＧ | ＨG x ｈＧ |  | 210,714 | 0 |
| MＴ | ＨT x ｈＴ |  | 1,120,402 | 2,317,958 |
| Mo | MＤ＋ＭＧ＋ＭＴ |  | 11,683,108 | 14,510,676 |
| Mg | Ｗg x ｈＧ |  | 23,035,062 | 10,290,748 |
| T | Thickness | mm | 30.0 | 40.0 |
| t0 | Design thickness | mm | 25.5 | 33.0 |
| h |  |  | 8.0 | 10.0 |
| h0 | sqrt（Bg0） |  | 124.2 | 185.9 |
| h/h0 |  |  | 0.06 | 0.05 |
| g１ |  |  | 7.00 | 16.00 |
| G1/g0 |  |  | 1.00 | 1.00 |
| F |  |  | 1.00 | 1.00 |
| F |  |  | 0.91 | 0.91 |
| E | F/h0 |  | 0.01 | 0.00 |
| T |  |  | 1.90 | 1.90 |
| U |  |  | 52.00 | 23.00 |
| V |  |  | 0.55 | 0.55 |
| D | U/V x h0 x g0^2 |  | 575,429 | 1,990,178 |
| L | (t0e+1)/T + t0^3 / d |  | 0.65 | 0.63 |
| Ｈ | Stress along hub spindle | N/mm2 | 166 | 42 |
| 1.5a - H | Never negative | N/mm2 | 7 | 131 |
|  |  |  | ＯＫ | ＯＫ |
| R | Radial direction stress | N/mm2 | 31 | 8 |
| a - R | Never negative | N/mm2 | 84 | 107 |
|  |  |  | OK | OK |
| Z |  | N/mm2 | 26.00 | 2.20 |
| T | Hoop stress | N/mm2 | -43 | 62 |
| a - T | Never negative | N/mm2 | 158 | 53 |
|  |  |  | OK | OK |
| (H + R)/2 |  | N/mm2 | 98.13 | 25.07 |
| a - (H + R) / 2 | Never negative | N/mm2 | 16.87 | 89.93 |
|  |  |  | OK | OK |
| (H + T)/2 |  | N/mm2 | -6.25 | 35.20 |
| a - (H+R) / 2 | Never negative | N/mm2 | 121.25 | 79.80 |
|  |  |  | OK | OK |

Table 9  Results of stress calculations for the flanges.

## Finite Element Analysis Results

Finite element analysis is performed based on the design described above to check the maximum deformation and stress distribution. The results are shown in Fig. 2 and Fig. 3. Maximum deformation (5.3mm) occurs on the outer-vessel inner wall (0.5mmt steel) where the deformation is along the radial direction. Maximum stress appears along the welded line on the outer-vessel inner wall, but this is less than the allowed pressure of the material.

Fig. 2  Result of finite element analysis (deformation).

Fig. 3  Result of finite element analysis (stress distribution).

## Drawings

All the drawing can be found in the MEG web site. These drawings are only for the vessel. These do not include the followings; Liquid nitrogen cooling pipe, feedthrough connectors, vacuum pump system, any port for sensors/gauges, gas inlet/outlet, drain for liquid phase purification. These must be included into the design before  construction.

## Detector Components

Followings are detector components currently considered to be used in the cryostat although they can be substituted by any kind of better components available now. So please consider each of them as one of the candidates that fulfill the minimum requirement.

### U-tight seal

A special kind of metal gasket is required for the inner vessel covers because they face liquid xenon whose temperature is –108 degree C. Currently “U-tight seal4” is considered to be used for the cryostat. Followings are the specification of it. The shape of the groove on the flange is determined by recommendation for using this metal gasket.

- Usage condition

  - Pressure: Vacuum - 0.3 MPa

  - Temperature: –110 ~ 100 degree C

  - Fluid: Liquid xenon

  - Frange and bolt: SUS316L

- U-tight seal dimension

  - Cross section: 5.5 mm diam

  - Material: Aluminum(outer), Stainless(Inner), Spring(Inconel)

### Vacuum pumping system

Two vacuum pumping system are considered to be used for the cryostat. One is for evacuating the inner vessel before filling xenon and the other is for the insulation vacuum layer which has to be evacuated continuously during the operation.

  - ALCATEL ATP400 for evacuating the inner vessel. 400 liter ceramic bearing turbo pump. The inlet flange type is DN 160 CF.

  - PFEIFFER TMU 261 for evacuating the outer vessel. 210 liter turbo pump. The inlet flange type is DN100 CF.

Both will be connected to the cryostat with gate valves.

### Refrigerator/Liquid nitrogen cooling pipe

A pulse tube refrigerator is equipped on the central nozzle on the top of the cryostat. The refrigerator works for pre-cooling the cryostat, liquefaction and re-condensation of xenon. The flange shape of the refrigerator is the same one used in the large prototype. Part of the drawing of the large prototype top cover is shown in Fig. 4 as a reference.

Fig. 4  Part of the drawing of the large prototype. The pulse tube refrigerator is installed in the 100mm diameter hole and fixed on the cover with twelve M12 bolts. Depth of the screw hole is 15mm. The groove for the metal gasket (helicoflex) is V120.

When cooling power of the refrigerator is not sufficient in such cases like fast liquefaction, power failure, or insulation vacuum break, liquid nitrogen cooling system will help to compensate the additional heat load. For this purpose a liquid nitrogen cooling pipe will be equipped on the cover of the central top nozzle as was done in the large prototype. A valve to control the nitrogen flow should be opened/closed automatically when needed. The system operating method is established in large prototype studies and proved to work in a safe manner.

Fig. 5  Part of the drawing of the large prototype to show how the liquid nitrogen cooling pipe is equipped. Details can be found in its drawing1.

### Feedthrough connectors

We will use about 800 PMTs in the cryostat. Two coaxial cables of HV and signal are needed for each of them. Requirements to the feedthrough are different between signal and HV. Feedthrough connectors for HV supply must be tolerant, at least, up to 2.0kV but do not need to transmit fast signal while for signal transmission of fast PMT signal is truly important. In Fig. 6 and Fig. 7 feedthrough connectors currently used in the large prototype are shown. The HV feedtrough is mounted on the top cover and fixed by cramps with o-rings between the flange and cover. The signal feedthrough is welded on a steal cover with is mounted on a flange. When constructing the cryostat for the final photon detector, we should pay attention to reduce difficult welding part in order to avoid unexpected leakage.

Fig. 6  HV feedthrough connector used in the large prototype. 260 wires in total.

Fig. 7  Signal feedthrough connector used in the large prototype. 256 channel in total.

### Sensors

Sensors to measure environmental variables like temperature and pressure in the cryostat will be located inside (temperature sensors, surface level meter) or mounted on the cryostat (pressure and vacuum gauges).

For reading out the sensors located in the cryostat we need feedthrough connectors although the number of channel is not quite large. Currently we are considering to use a commercially available feedthrogh connectors mounted on CF flanges. Flange type for the pressure and vacuum gauges will be also CF type.

For mounting them we have to prepare several ports on the top cover with CF flanges. Details have not been determined yet. A list of sensors is shown in Table 10.

### Purification system

In this design liquid phase purification is not considered. Only gas phase purification is taken into account. For implementing liquid phase purification we need a drain at the bottom of the cryostat and liquid xenon inlet port on the top. These must be newly added into the design.

## Flow sheet of the system

Whole system is schematically shown in . However as mentioned above, liquid phase purification scheme is not included into this figure. This must be included in future.

### Sensors

| Section | Discription | Control | Mark |  | Sensor | Range |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Photon Detector | Inner Vessel TEMP(Bot) |  | TI-101 |  | PT100 | 77K~327K |  |
|  | Inner Vessel TEMP(Middle) |  | TI-102 |  | PT100 | 77K~327K |  |
|  | Inner Vessel TEMP(Top) |  | TI-103 |  | PT100 | 77K~327K |  |
|  | Vac. Vessel TEMP(Top) |  | TI-104 |  | PT100 | 77K~327K |  |
|  | LN2 Cooling Pipe TEMP(In) |  | TI-105 |  | PT100 | 77K~327K |  |
|  | LN2 Cooling Pipe TEMP(Out) | O | TIC-106 |  | PT100 | 77K~327K |  |
|  | LN2 Pipe VAC Vessel TEMP(In) |  | TI-107 |  | PT100 | 77K~327K |  |
|  | LN2 Pipe VAC Vessel TEMP(Out) | O | TIC-108 |  | PT100 | 77K~327K |  |
|  | LXe Level | O | LIA-101 |  |  |  |  |
|  | LXe Pressure | O | PI-101 | PT-101 |  | -0.1~0.5MPa |  |
|  | Inner Vessel VAC Gauge |  | PI-102 |  |  | 10^-5~0.1MPa |  |
|  | Vac. Vessel VAC Gauge |  | PI-110 |  |  | 10^-5~0.1MPa |  |
|  | Vac. Vessel Pressure |  | PI-112 | PT-112 |  | -0.1~0.5Mpa |  |
| Refrigerator/ | Cold Head TEMP | O | TIC-201 |  | PT100 | 77K~327K |  |
| Compressor | Top Flange TEMP |  | TI-202 |  | PT100 | 77K~327K |  |
|  | Heater Power W(VxI) |  |  |  |  |  |  |
|  | Compressor Delivery Pressure |  | PI-211 | PT-211 |  | 0~3.5MPa |  |
|  | Compressor Suction Pressure |  | PI-212 | PT-212 |  | 0~3.5MPa |  |
|  | Cooling Water Flow Meter |  | FIA-211 |  |  | ~10L/min |  |
|  | Cooling Water TEMP(In) |  | TIA-211 |  |  | -20~50 degreeC |  |
|  | Cooling Water TEMP(Out) |  | TI-212 |  |  | -20~50 degreeC |  |
|  | On/Off Status |  |  |  |  |  |  |
|  | Interlock INFO(Overload) |  |  |  |  |  |  |
|  | Interlock INFO(TEMP) |  |  |  |  |  |  |
|  | Interlock INFO(Pressure) |  |  |  |  |  |  |
| Purification | Getter Purifier Pressure (In) |  | PI-301 | PT-301 |  | -0.1~0.5MPa |  |
| (Gas Phase) | Getter TEMP |  | TI-301 |  |  | 0~500 degreeC |  |
|  | Getter Status (Stanby/Purify) |  |  |  |  |  | not shown in the diagram |
|  | Circulator Pump On/Off Status |  |  |  |  |  |  |
| Storage Tank | #1 Vessel TEMP |  | TI-401 |  | PT100 | 77K~327K |  |
|  | #2 Vessel TEMP |  | TI-402 |  | PT100 | 77K~327K |  |
|  | #3 Vessel TEMP |  | TI-403 |  | PT100 | 77K~327K |  |
|  | #4 Vessel TEMP |  | TI-404 |  | PT100 | 77K~327K |  |
|  | #5 Vessel TEMP |  | TI-405 |  | PT100 | 77K~327K |  |
|  | #6 Vessel TEMP |  | TI-406 |  | PT100 | 77K~327K |  |
|  | #7 Vessel TEMP |  | TI-407 |  | PT100 | 77K~327K |  |
|  | #8 Vessel TEMP |  | TI-408 |  | PT100 | 77K~327K |  |
|  | #1 Vessel Pressure |  | PI-401 | PT-401 |  | -0.1~10MPa |  |
|  | #2 Vessel Pressure |  | PI-402 | PT-402 |  | -0.1~10MPa |  |
|  | #3 Vessel Pressure |  | PI-403 | PT-403 |  | -0.1~10MPa |  |
|  | #4 Vessel Pressure |  | PI-404 | PT-404 |  | -0.1~10MPa |  |
|  | #5 Vessel Pressure |  | PI-405 | PT-405 |  | -0.1~10MPa |  |
|  | #6 Vessel Pressure |  | PI-406 | PT-406 |  | -0.1~10MPa |  |
|  | #7 Vessel Pressure |  | PI-407 | PT-407 |  | -0.1~10MPa |  |
|  | #8 Vessel Pressure |  | PI-408 | PT-408 |  | -0.1~10MPa |  |
|  | Liquefy Line Pressure | O | PI-421 | PT-421 |  | -0.1~10MPa |  |
|  | Liquefy Regulator TEMP | O | TIC-421 |  | PT100 | 77K~327K |  |
|  | Liquefy Line Flow Meter |  | FI-421 |  |  | 0~60L/min |  |
|  | Recovery Flow Meter |  | FIQ-435 |  |  | 0~60L/min |  |
|  | UNIT A LN2 Level | O | LIA-431 |  |  | 0~FULL |  |
|  | UNIT A Weight |  |  |  |  | 0~2000kg | not shown in the diagram |
|  | UNIT B LN2 Level | O | LIA-432 |  |  | 0~FULL |  |
|  | UNIT B Weight |  |  |  |  | 0~2000kg | not shown in the diagram |
|  | UNIT C LN2 Level | O | LIA-433 |  |  | 0~FULL |  |
|  | UNIT C Weight |  |  |  |  | 0~2000kg | not shown in the diagram |
|  | UNIT D LN2 Level | O | LIA-434 |  |  | 0~FULL |  |
|  | UNIT D Weight |  |  |  |  | 0~2000kg | not shown in the diagram |
| LN2 System | #1 LN2 Dewer Pressure |  | PI-501 |  |  | 0~1Mpa |  |
|  | #1 LN2 Dewer Level |  | LIA-501 |  |  | 0~FULL |  |
|  | #2 LN2 Dewer Pressure |  | PI-502 |  |  | 0~1Mpa |  |
|  | #2 LN2 Dewer Level |  | LIA-502 |  |  | 0~FULL |  |
|  | #3 LN2 Dewer Pressure |  | PI-503 |  |  | 0~1Mpa |  |
|  | #3 LN2 Dewer Level |  | LIA-503 |  |  | 0~FULL |  |

Table 10  List of sensors planned to be used.

## Pipe dimension for xenon and nitrogen transfer

Pipe dimension for xenon and nitrogen transfer was investigated. Results are summarized in Table 11 and pressure drop calculation result is summarized in Table 12 for two typical xenon flow rates.

|  |  | Material | Size | Flow rate | Gas | Press | Temp | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | MPa | degree C |  |
| 1 | xenon gas supply/recovery | SUS304 | 12.7 diam. x t1 | 6Nm3/hr | GXe | 0.1 | 30 | 34mm diam. |
| 2 | Purified gas | SUS304 | 9.53 diam. x t1 | 1.5Nm3/hr | GXe | 0.15 | 30 |  |
| 3 | gas recovery from the vacuum layer | SUS304 | 25A sch5S | 30Nm3/hr | GXe | 0 | 30 |  |
| 4 | liquid xenon recovery | SUS304 | 9.53 diam. x t1 | 52L/hr | LXe | 0.1 | -108~30 |  |
| 5 | LN2 supply (cooling) | SUS304 | 9.53 diam. x t1 | 10L/hr | LN2 | 0.5 | -196 |  |
| 6 | LN2 supply (tank) | SUS304 | 15A sch5S | 500L/hr | LN2 | 0.5 | -196 | 21.7mm diam. |

Table 11  Dimension of pipes for liquid nitrogen and xenon transfer.

| Pressure Drop Calculation | Unit |  |  |
| --- | --- | --- | --- |
| Gas |  | Xenon | xenon |
| flow rate | Nm3/Hr | 3 | 6 |
| Temperature | degree C | 25 | 25 |
| Pressure Drop Calculation | kg/cm2G | 1 | 1 |
| outer dimension of the pipe | mm | 9.53 | 12.7 |
| thickness of the pipe | mm | 1 | 1 |
| density of the gas | kg/Nm3 | 5.391 | 5.391 |
| Viscosity | kg/ms | 0.0000231 | 0.0000231 |
| length of the pipe | mm | 10 | 10 |
| number of elbows |  | 10 | 2 |
| number of Ts |  | 2 | 2 |
| number of valves |  | 2 | 2 |
| Pressure drop (smooth pipe) | kg/cm2 | 0.296 | 0.224 |
| Pressure drop (rough pipe) | kg/cm2 | 0.678 | 0.498 |
| Flow speed | m/sec | 10.4 | 10.3 |
| Inner diameter | m | 0.00753 | 0.0107 |
| Density | kg/m3 | 9.719698016 | 9.719698 |
| corresponding length of the pipe | m | 18.1324 | 21.556 |
| Reynolds |  | 32884.50041 | 46284.166 |
| friction coefficient (smooth pipe) |  | 0.005750842 | 0.0053154 |
| friction coefficient (rough pipe) |  | 0.01318011 | 0.0118034 |

Table 12  Estimation of pressure drop in xenon transfer pipe for two typical flow rate.

## References

# Footnotes

Taiyo Toyo Sanso Co., Ltd.  http://www.saan.co.jp/english/index.html

# Endnotes

Drawing of the large prototype can be found at http://meg.web.psi.ch/subprojects/calorimeter/prototype.html

Photograph of the cooling pipe in the large prototype cryostat can be found at http://meg.web.psi.ch/subprojects/calorimeter/photos/vessel/tn/photo10.jpg.html

http://www.usui.co.jp/eng/products/12_u-tight_seal.html

http://meg.web.psi.ch/subprojects/calorimeter/drawings/PD_vessel.html

http://www.lesker.com/cfdocs/newweb/Vacuum_Pumps/Individual_Pumps/Turbo_and_Drag_Pumps/TurboMolecular_CeramicBearing_Alcatel_ATP400_ATP900.cfm

http://www.pfeiffer-vacuum.de/shop/frontend/product.php3?productID=PMP02825